I	PURPOSE

The purpose of this procedure is to provide a system and instructions and to assign responsibilities to identify calibrate, and maintain measuring and test equipment that meet the requirements of ANSI-Z540.

II	APPLICATION

This procedure applies to measuring and test equipment, to comparative reference hardware (such as templates) and for control of production processes. This procedure concerns Quality Assurance and Production.

III	PROCEDURE

1.	Measurement Identification and Selection of Equipment

1.1	Identification of measurements to be made and the tolerance of the measured characteristics are provided in customer-supplied drawings and specifications and/or control plans or industry accepted references.

1.2	Gauges, instruments, and other measuring and testing equipment are selected on the basis of their capability to provide the necessary accuracy of the measurement. Production personnel are responsible for the selection of the appropriate measuring and test equipment.

2.	Calibration

  - Quality Assurance is responsible for maintenance and control of all inspection, measuring and test equipment, including equipment belonging to employees or on loan.  Equipment is to be calibrated by an approved vendor, or in accordance with written instructions, unless calibration is simple and obvious. Applicable instructions may specify that each type of equipment has acceptable limits of temperature, pressure, humidity, and other environmental conditions that may affect calibration.

  - ManTech does not maintain facilities to calibrate measuring and test and inspection equipment.  The equipment is subcontracted to a facility meeting, the requirements of ANSI Z540 and audited by an independent entity for compliance. Quality Assurance personnel verify that purchase orders issued for the calibration of measuring, test and inspection equipment meet the standards of ANSI Z540.

2.3	Calibration of measuring and test equipment is carried out using calibration instruments or standards certified to have a known relationship to a nationally recognized standard.  This relationship is identified on the calibration record.  Equipment that is sent out for calibration is required to be returned with certificates that are likewise traceable to national standards.  Quality Assurance maintains calibration records and certificates.

2.4	Calibrated equipment is labeled with a sticker indicating the date of calibration, identification markings and the due date for next calibration.  Equipment with a past-due calibration date or without a calibration sticker is not used and is to be returned to Quality Assurance.

2.5	Quality Assurance maintains a list of all active measuring and test equipment, whether or not the company owns the equipment.  The list identifies each piece of equipment by its name, type, size, serial number, calibration frequency, last calibration date, and next calibration due date.  The list is verified at least once a month.

3.	Product-Specific Comparative Reference Hardware

3.1	Jigs, fixtures, templates, patterns, and samples used in production and inspection are uniquely identified with a unique identification marking.

3.2	The comparative reference hardware where applicable, is entered on the list of active measuring and test equipment (see Subsection 2.5), and is periodically checked for accuracy.  Records of the accuracy checks are maintained.

4.	Storage and Maintenance

4.1	Measuring and test equipment shall be stored and issued through the tool room.  The equipment is maintained, stored, and handled in such a way as to preserve its accuracy and fitness for use.  Equipment that is out of calibration or is otherwise not fit for use is withdrawn from the inspection and production areas, is segregated or forwarded to the QA department for repairs and recalibration or disposal.

5.	Nonconforming Equipment

5.1	When a piece of measuring or test equipment is found to be out of calibration or appears to give inaccurate readings, the accuracy is verified.  If the inaccuracy is confirmed, that the equipment is verified out of calibration and/or the readings are outside of required accuracy, Quality Assurance investigates and assesses the validity of measurements for which the equipment was previously used.  Identification of equipment and the impact of its use on acceptance of products are reported in a nonconformance report, in accordance with SOP-4P-05, Control of Nonconforming Product. The customer is notified if suspect material has been shipped.

6.	Equipment Exempted From Calibration

6.1	Inspection and test equipment may be exempted from calibration when used in situations where accuracy of measurement is not important, or where the measurement does not have any relation to product verification or process control. Such equipment is labeled with stickers warning that it is not calibrated.  Production and inspection personnel are to be made aware of the limitations in using uncalibrated equipment.

IV	Associated Documents

SOP–Q-03, Quality Records

SOP–Q-04, Receiving Inspection

SOP–P-05, Control of Nonconforming Product